## List1
|  |  | BEFORE RISK MITTIGATION | [6 empty] | AFTER RISK MITTIGATION
MEDICAL DEVICE CHARACTERISTIC NUMBER | DETERMINATION OF THE CHARACTERISTICS OF A MEDICAL DEVICE | APPLIES TO THE MEDICAL DEVICE | RISK IDENTIFICATION | RISK DESCRIPTION | ESTIMATION OF THE PROBABILITY OF RISK OCCURENCE (V) | ESTIMATION OF RISK SEVERITY LEVEL (Z) | RISK MITIGATION | RISK ACCEPTABILITY / RESIDUAL RISK / Clinical benefit | EVIDENCE OF RISK MITIGATION | ESTIMATION OF THE PROBABILITY OF RISK OCCURENCE (V) | ESTIMATION OF RISK SEVERITY LEVEL (Z) | RISK MITIGATION | ACCEPTABILITY OF THE RISK / RESIDUAL RISK / Clinical benefit | EVIDENCE OF RISK CONTROL
1. | What is the intended use and how should the medical device be used? | Yes | R 1.1 | Sharp edges or corners: In case of sharp edges of individual mechanical parts of the device, the operator may be injured. | V2 | Z2 | The components that can be handled by medical personnel are designed with chamfered edges or with a radius by all sharp edges. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Drawing documentation | V1 | Z3 | The sharp edges of the individual parts of the device are inspected during the receiving inspection and final check of the manuafacturer. Sharp edges are eliminated by geometry and prescribed chamfering / rounding on drawings. Compliance with these regulations is checked by receiving inspection, so a dangerous situation is prevented. | The risk of injury to the patient or operator from sharp edges is due to:- carrying out inspections, - prescribing a requirement for drawing documentation, acceptable. The clinical benefit outweighs the risk of sharp edge injury | Drawing documentation / input control
|  |  | R 1.2 | Use in the pediatric population and the adult population weighing less than 50 kg | V3 | Z3 | According to the ventilator specification, the ventilator can only be used for adult patients weighing more than 50 kg. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Instructions for use. | V2 | Z3 | Indication of the recommended weight of the ventilated patient in the Instructions for use. | The risk of ventilation for a patient weighing less than 50 kg or for a child is acceptable with respectto the warning in the instructions for use and the placement of a warning label at the breathing circuit connection. | Instructions for use.
|  |  | R 1.3 | Ventilation function failure. | V2 | Z4 | A manual resuscitation bag must always be available close to the ventilator | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Instructions for use. | V2 | Z4 | Alarms indicating failure of the ventilation function. | A manual resurrection bag is a standard emergency equipment for every intensive care unit / gurney. The risk is acceptable, the requirement for the presence of a manual resuscitation bag is also mentioned in the Instructions for use. | Instructions for use.
|  |  | R 1.4 | Application of positive overpressure to the patient's lungs above the tolerance limit. | V1 | Z4 | Pressure limitation on 2 stages. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical drawing and specifications of alarms. | V1 | Z4 | Dangerous overpressure level alarms are factory set and adjustable by the operator. | The risk of high airway pressure is limited both by design and by the user being alerted to this fact by alarms. | Instructions for use.
|  |  | R 1.5 | Loss of positive overpressure after expiration | V2 | Z4 | The ventilator is equipped with a system to prevent a drop in positive pressure after expiration by generating so-called PEEP. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical drawing and specifications of alarms. | V1 | Z4 | EEP on the manometer located on the housing of the device. | The risk of losing positive positive pressure after expiration is limited in construction | Instructions for use.
2. | Is the medical device intended for implantation? | No
3. | Is the medical device intended to come into contact with the body of the patient or other persons? | Yes | R3.1 | Patient inhales small particles of dirt from the device. | V2 | Z3 | The inspirational part consists of biocompatible components - plastic tubes, electromechanical valves. These components undergo an output check at the manufacturer and these are tested after being installed into CoroVent. Any loose particles or impurities are entrained by the gas jet from the ventilator. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Final factory check report of CoroVent device. | V1 | Z2 | The ventilator is set to maximum flow during the final factory check, the output of the ventilator, ie the inspiratory part of the ventilation circuit is open to the environment and the gas flows for at least 3 minutes, thus removing all unwanted elements from the internal ventilator structures. | The risk of releasing unwanted elements is acceptable due to final check performed. | Final check.
|  |  | R3.2 | Metal ventilator housing - leakage current in contact with the patient | V2 | Z3 | Leakage current measurement during CoroVent output control | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Final factory check. | V2 | Z3 | Leakage current measurement during CoroVent final factory check. The value of leaked current must meet the standard. | Acceptable risk | Final check. Measurement of electrical safety as part of periodical technical safety inspection
|  |  | R3.3 | Electric current shock from pressure sensors in the patient circuit | V2 | Z3 | The risk is eliminated by the used 24V source, which is a safe voltage | The risk was minimized by using a separate power source. | Technical documentation for the power supply source. | V2 | Z2 | The risk is eliminated by the used 24V source, which is a safe voltage. | The risk was minimized by using a separate source and using a device housing. | Receiving inspection.
4. | Which materials or components are used in the device or are in contact with it? | Yes | R4.1 | The materials used for the distribution of the gaseous ventilation mixture are not applicable for the distribution of high oxygen concentrations. | V1 | Z3 | Only components safe for use with high or 100% oxygen concentration are used in the oxygen inspiratory branch. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical documentation of used components, final factory check. | V1 | Z2 | The type of components used for internal distribution will be clearly stated in the list of components used. The material used will be suitable for oxygen distribution. The material does not come into contact with any grease during assembly. | The risk of using unsuitable materials is acceptable due to the performed receiving inspection of used materials and final factory check. | Receiving inspection of delivered and used components. Final factory check.
|  |  | R4.2 | The housing material of the device is not compatible with disinfectants used in hospital facilities. | V1 | Z2 | The housing material is compatible with strong disinfectants. | The risk is acceptable, material degradation is generally a slow process. | Technical documentation of used components, final factory check. Recommended definitions are listed in the Instructions for use. | V1 | Z2 | The risks and recommended disinfectants are listed in the Instructions for use. | The risk is acceptable. | Receiving inspection of delivered and used components. Final check. Instructions for use.
5. | Is energy supplied or withdrawn to the patient? | Yes | R5.1 | Applying excessive pressure to the patient's lungs (see R1.4) | V1 | Z4 | Pressure limitation on 2 stages by means of pressure reducing valves. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical drawing and specifications of alarms. | V1 | Z4 | Dangerous overpressure level alarms are factory set and adjustable by the operator. | The risk of high airway pressure is limited both by design and by the user being alerted to this fact by alarms. | Final check protocol.
|  |  | R5.2 | Cool the patient by delivering a cold breathing mixture to the patient's breathing circuit. | V2 | Z4 | Use of HME (Heat and Moisture Exchanger) within the ventilation circuit. | The risk is acceptable, the ongoing process is slow from a clinical point of view. HME is a standard consumable designed to store heat and moisture in the patient circuit breathing mixture. | Technical documentation of used components, Instructions for use and operator training. | V2 | Z3 | Emphasis on the description of the recommended HME filter and their regular replacement according to the Instructions for Use. | Based on the implemented measures, the frequency of risk is reduced. | Instructions for use, operator training.
6. | Are any substances given or removed from the patient? | No
7. | Are biological materials processed by a medical device for their subsequent re-use, transfusion or transplantation? | No
8. | Is the medical device supplied as sterile, or is it intended for user sterilization, or are other microbiological procedures applicable? | No
9. | Is the medical device intended to be routinely cleaned and disinfected by the user? | Yes | R9.1 | Use of inappropriate disinfection of all surfaces. | V3 | Z2 | The housing material is compatible with strong disinfectants. | The risk is acceptable, the chemical events associated with the side effects of disinfectants are slow, but information labels with important information can be damaged. More frequent Periodical technical safety inspection of the device is desirable to detect and eliminate defects and deficiencies in the device. | Instructions for use. Periodical technical safety inspection - every 3 months. | V2 | Z2 | Reffer to Instructions for use for instructions for cleaning, maintenance and use of the recommended cleaning and disinfecting agents. | Operators of the device are informed about the function of individual elements. More frequent inspections will reveal damaged information labels or other deficiencies. As part of the standard maintenance, the maintenance technician will carry out the repair on site. | Instructions for use.
|  |  | R9.2 | Increased antiviral disinfection associated with the occurrence of SARS-CoV-2. | V3 | Z2 | The housing material is compatible with strong disinfectants and can withstand frequent disinfection processes. | The risk is acceptable, the chemical events associated with the side effects of disinfectants are slow, but information labels with important information can be damaged. More frequent Periodical technical safety inspection of the device is desirable to detect and eliminate defects and deficiencies in the device. | Instructions for use. | V2 | Z2 | Reffer to Instructions for use for instructions for cleaning, maintenance and use of the recommended cleaning and disinfecting agents. | Operators of the device are informed about the function of individual elements. More frequent inspections will reveal damaged information labels or other deficiencies. As part of the standard maintenance, the maintenance technician will carry out the repair on site. | Instructions for use. Periodical technical safety inspection - every 3 months.
|  |  | R9.3 | Leakage of disinfectants into the device and its subsequent failure or electric shock | V3 | Z3 | The CoroVent housing used is class IP21, the cover is resistant to fluid leakage. | The risk of disinfection leaking is unacceptable. Internal circuits, sensors and active elements - valves - can be damaged due to leakage. The operation of the device can then be adversely affected. | Taken into account in technical documentation, Instructions for use and operator training, including supporting personnel. | V3 | Z3 | The documentation will state the specification of the used housing of the device, the instructions for cleaning and disinfection of the device surfaces will be given in the Instructions for use, and the operator will be acquainted with proper and safe procedures as part of the operator's training. | The risk is acceptable. | Drawing documentation. Instruction for use. Operator training.
10. | Is the medical device designed to change the patient's environment? | Yes | R10.1 | Enrichment of inhaled air with oxygen - poor mixing of the mixture | V2 | Z3 | Treated by monitoring the pressures in the inspiratory branches, testing the formation of mixtures during production, the service life of the used elements - valves - is known. | The risk is acceptable in adult patients, the higher oxygen content in the ventilation mixture represents an acceptable risk for the patient, the risk of fire increases. | Technical documentation, maintenance documentation - prescribed service life of the components used, recommendations in the Instructions for use for the necessary checks of patients' vital functions by means of external vital functions monitoring. | V2 | Z3 | Technical documentation, maintenance documentation - prescribed service life of the components used and the necessary observance of maitenance intervals, recommendations in the CoroVent Instructions for use foe the necessary control of patients' vital functions by means of external vital functions monitoring. | The risk is acceptable for adult patients, the higher oxygen content in the ventilation mixture represents an acceptable risk for the patient, the risk in of fire increases. | Technical documentation, maintenance documentation - prescribed service life time of used components, recommendations in the Instructions for use for the necessary control of patients' vital functions by means of external vital functions monitoring. | Technical documentation, maintenance documentation - prescribed service life of the components used, recommendations in the Instructions for use for the necessary checks of patients' vital functions by means of external vital functions monitoring
|  |  | R10.2 | Drying of the patient's mucous membranes due to poor or no humidification of the ventilation gas mixture. | V3 | Z3 | An HME filter must always be used in the ventilation circuit, regular replacements must be provided. | The risk is unacceptable, it represents damage to the mucouses of the respiratory tract. | Instructions for use, operator training. | V2 | Z3 | Emphasis on the description of the ventilation circuit and use only the ventilation circuits recommended by the manufacturer. | Based on the implemented measures, the frequency of risk is reduced. | Instructions for use, operator training.
|  |  | R10.3 | Respiratory temperature - heating or cooling of the respiratory tract (see R5.2 and R5.3) | V3 | Z3 | Temperature measurement of inspired gas during testing, ventilation, HME filters. | When using HME, the risk is minimal. | V Návodu k použití popsány doporučené HME filtry. | V2 | Z3 | Emphasis on the description of the recommended HME filter and their regular replacement according to the Instructions for use. | Based on the implemented measures, the frequency of risk is reduced. | Instructions for use, operator training.
|  |  | R10.4 | Excessive pressure changes in the airways and inside the chest. | V3 | Z4 | Introduction of alarms of pressure values, safety pressure relief valve preventing the supply of excessive pressure to the patient's lungs, control of the pressure at the entrance to the patient's airways using a manometer. | The risk is unacceptable, it poses a risk of damage to the patient's lungs. | The recommended HME filters are described in the Instructions for Use. | V2 | Z3 | Emphasis on the description of safe adjustment of all pressure and volume elements of ventilation and control of their parameters on the display and available manometers. Emphasis on individual setting of pressure and volume alarms for each patient before the start of ventilation and their continuous control during ventilation. | Based on the implemented measures, the frequency of risk is reduced. | Instructions for use, operator training.
11. | Are measurements performed? | Yes | R11.1 | Inaccuracy of airway pressure measurement influencing the user's decision to change ventilation parameters | V3 | Z3 | 2 stages of airway pressure measurement - in the expiratory arm and at the outlet of the device into the airways. CoroQuant sensor The lowest sensitivity in the measuring range ± 10 l / min of gas is at a stable temperature in the range of ± 1.7 l / min. The whole measuring chain is further loaded with a measurement error of ± 0.5 l / min with a temperature change of ± 1 ° C, where there is a voltage drift of the zero flow point and balancing of the measurement of the integrated volume of the inspiration and expiration curve.)
12. | Does the medical device provide evaluable data? | Yes | R12.1 | Measured data - pressure parameters | V3 | Z3 | Absolute pressure sensors and pressure relief valve installed | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Pneumatic diagram. | V2 | Z3 | Alarm messages 23, 99, 2, 3, 10, 11 | The risk of incorrect pressure measurement has been reduced to an acceptable value | Instructions for use.
|  |  | R12.3 | Measurement of inlet gas pressure | V3 | Z3 | Working pressure sensors in inspiratory branches for oxygen and air | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Pneumatic diagram. | V2 | Z3 | Alarm messages 24, 99, 2, 3 | The risk of ventilator malfunction has been sufficiently reduced | Instructions for use.
13. | Is the medical device intended for use in conjunction with other medical devices, medicines or other medical techniques? | Yes |  | see R17.1 | V3 | Z2 | Use of a standard patient circuit with HME, which is a disposable consumable. | The risk is acceptable. | Instructions for use. | V2 | Z2 | Checking the operation of the device when commissioning the CoroVent ventilator. Use of a standard patient circuit with HME, which is a disposable consumable. | The risk is acceptable. | Instructions for use, operator training.
14. | Is there an unwanted release of energy or substances? | Yes | R14.1 | Due to losses, the control electronics release thermal energy | V3 | Z3 | The housing of the device contains two ventilation grilles. | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical documentation. | V2 | Z2 | The ventilator contains an active fan located inside the housing of the device. The fan ensures sufficient air circulation inside the CoroVent housing. | The risk has been sufficiently reduced. | See technical documentation.
|  |  | R14.2 | During operation, an electromagnetic field is generated which can affect the surrounding devices | V3 | V2 | The device is designed to minimize the generation of electromagnetic fields. EMC tests performed. | The risk is acceptable. | Technical documentation, Instructions for use, Test Report. | V3 | V2 | The Instructions for use contain information on the minimum distance of other devices in order to eliminate the interaction by the electromagnetic field. | The risk is acceptable. | Instructions for use, operator training.
|  |  | R14.3 | Electric shock to operators | V3 | Z3 | The CoroVent ventilator is designed to prevent electric shock. Electrical safety tests performed at EZU. | The risk is unacceptable. | Instructions for use, power source certificates, final factory check. | V2 | Z2 | A power supply for medical purposes is used to supply critical parts, the individual parts of the ventilator are supplied with a safe voltage of 24 V. After the completion of the device, the check of the connection of the electrical part is included and the parameters of the electrical safety are measured. It is recommended to operate the device only within VDO or ZIS. | The risk has been sufficiently reduced. | Used components, final check, Instructions for use
|  |  | R14.4 | Electric shock to patient | V3 | Z3 | The CoroVent ventilator is designed to prevent electric shock. Electrical safety tests performed at EZU. | The risk is unacceptable. | Instructions for use, power source certificates, final factory inspection. | V2 | Z2 | A power supply for medical purposes is used to supply critical parts, the individual parts of the ventilator are supplied with a safe voltage of 24 V. After the completion of the device, the check of the connection of the electrical part is included and the parameters of the electrical safety are measured. It is recommended to operate the device only within VDO or ZIS. | The risk has been sufficiently reduced. | Used components, final check, Instructions for use
|  |  | R14.5 | Heat injury when in contact with the metal housing of the device during its significant heating. | V2 | Z2 | Active cooling of the device by means of a fan. | Risk unacceptable due to the possibility of contact of the operator or patient with the metal housing of the device. | Technical documentation. | V2 | Z2 | Ensured air circulation inside the device using a fan on the back of the device and the possibility of air leakage through the grille on the side of the device. Circulating air ensures equal heat distribution. | The risk has been sufficiently reduced. | Technical documentation.
|  |  | R14.6 | Release of substances from the housing after contact with disinfectants. | V2 | Z2 | Housing material stainless steel for food industry 1.4301, thickness 1mm. The material does not release substances in contact with strong disinfectants. | The risk is acceptable in relation to the material used. | Technical documentation of material for the production of the device housing. | V1 | Z2 | The recommended instructions for cleaning and maintenance of the device housing are given in the Instructions for use. | The risk of releasing substances is reduced by using only recommended disinfectants. | Instructions for use.
15. | Is the medical device sensitive to the environment? | Yes | R15.1 | Due to the temperature, the measured value on the pressure sensor may change | V3 | Z3 | Selection of quality pressure sensors with time stability. | The risk is unacceptable due to the need for accurate pressure measurements. | Technical documentation for testing of pressure sensors. |  |  | The temperature drift of the pressure sensors is compensated electronically. | Risk reduced by operation in a temperature-stable room. | Instructions for use, operator training.
|  |  | R15.2 | Due to the temperature, the device may overheat (also 14.1) | V3 | Z3 | Ventilation grilles (measuring the temperature of inspired gas during testing) | The risk is unacceptable and the risk needs to be further mitigated. | Technical documentation. | V2 | Z3 | Active fan cooling | The risk has been sufficiently reduced. | Technical documentation.
|  |  | R15.3 | Due to the high pressure at the inlet of the device, the parameters of the supplied mixture may change | V3 | Z3 | The inlet pressures of air and oxygen are reduced to an operating pressure of 2 bar at an outlet flow of 50 L / min. Working pressures are sensed by pressure sensors. A drop in operating pressures below the tolerable limit will be triggered by the Insufficient Oxygen Source Pressure (E2) or Insufficient Air Source Pressure (E3) alarm. Unbearable difference between working pressures of oxygen and air triggers alarm large difference of working pressures of oxygen and air - service (E99) | The risk is acceptable, a higher concentration of O2 does not have a significant effect on the critical influence of the patient's vital functions, a low concentration of O2 will be reflected in the standard measured parameter SPO2 - monitored by the operator. | Technical documentation. | V3 | Z2 | Pressure reducing valves unctionality during the final factory check of the device. | The risk of changing the parameters of the supplied ventilation mixture is acceptable due to the used reduction valves and due to the control of their correct function by an independent pressure sensor. | Technical documentation. Instructions for use. Periodical technical safety inspection - every 3 months.
|  |  | R15.4 | Sensitivity to fluctuations in electricity supply | V3 | Z3 | Use of a quality stabilized source. | The risk is undesirable. Necessity to use a backup battery source with surge protection. | Technical documentation. | V2 | Z2 | Installation of a battery backup power supply with a current fuse. | The risk has been sufficiently reduced. | Technical documentation, final inspection
|  |  | R15.5 | Sensitivity to pressure fluctuations in the supply of medical gases | V3 | Z3 | The inlet pressures of air and oxygen are reduced to an operating pressure of 2 bar at an outlet flow of 50 L / min. Working pressures are sensed by pressure sensors. A drop in operating pressures below the tolerable limit will be triggered by the Insufficient Oxygen Source Pressure (E2) or Insufficient Air Source Pressure (E3) alarm. Unbearable difference between working pressures of oxygen and air triggers alarm large difference of working pressures of oxygen and air - service (E99) | The risk is acceptable. | Technical documentation. | V3 |  | Working pressures are sensed by pressure sensors. A drop in operating pressures below the tolerable limit will be triggered by the Insufficient Oxygen Source Pressure (E2) or Insufficient Air Source Pressure (E3) alarm. The unbearable difference between the working pressures of oxygen and air triggers the alarm. acid pressures. and air - service (E99) | Due to the used pressure reducing valves capable of covering fluctuations in the permitted range of pressures in medical gas distribution and subsequent control of working pressures by independent pressure sensors, based on which an alarm will be equipped in the event of a non-standard situation, the risk is acceptable. | Technical documentation, Instructions for use.
16. | Does the medical device affect the environment? | Yes | R16.1 | Noise load in the environment due to noisy ventilator operation | V5 | Z2 | Noise measurement below 60 dB | Acceptable risk compared to clinical benefit | Final factory check report. | V5 | Z2 | Acceptable risk compared to clinical benefit. | Acceptable risk compared to clinical benefit | Instructions for use.
|  |  | R16.2 | Noise pollution due to alarms | V5 | Z2 | The volume and tone of the alarm is according to the standards | The risk is acceptable. | Final factory check report. | V5 | Z2 | Acceptable risk compared to clinical benefit. | Acceptable risk compared to clinical benefit | Instructions for use.
|  |  | R16.3 | Genesis of the electromagnetic field when the device is running | V5 | Z2 | The genesis of the electromagnetic field is minimized by the design of the device. EMC tests will be performed on EZU. | The risk is acceptable. | Instructions for use. EZU test report. | V3 | V2 | The Instructions for use contain information on the minimum distance of other devices in order to eliminate the interaction by the electromagnetic field. | The risk is acceptable. | Instructions for use, operator training.
17. | Are significant consumables or accessories used with the medical device? | Yes | R17.1 | Unsuitable patient ventilation circuits | V3 | Z3 | Use only patient ventilation circuits without active humidification. | The risk is acceptable. | Instructions for use. | V2 | Z3 | The Instructions for use contain information on possible risks from other than recommended ventilation circuits. | The risk is acceptable. | Instructions for use, operator training.
18. | Is maintenance or calibration required? | Yes | R18.1 | Incorrect resetting of the CoroQuant flow sensor | V3 | Z3 | Zeroing the sensor before starting the patient's ventilation or as needed by the operator. | The risk is addressed in the Instructions for use. | Instructions for use. | V3 | Z3 | Resetting the sensor is always recommended after installing a new flow sensor, and always in case of suspected malfunction. | The risk is acceptable. | Instructions for use, operator training.
19. | Is the software part of the medical device? | Yes | R19.1 | Function failure due to a program error | V2 | V4 | Performing software tests during and after software production." "Unit tests: Testing of individual units (POU), always after the completion of writing each unit Integration tests: After connecting individual units into two units (Visualization and PLC) at the interface level Ie testing whether the Visualization interacts correctly with the PLC module interface Acceptance tests: After compiling the whole project, in the presence of people experienced in the field of lung ventilation. | The risk of ventilator failure is unacceptable compared to the clinical benefit. | Final factory check, 8 hours of trial run on artificial lungs - long-term test, internal tests ... | V2 | V3 | The stability of the SW is performed within the tests after the completion of the CoroVent device. In case of failure of the SW of the ventilation system, it is necessary to have in the vicinity of the ventilator, resp. patient hand resuscitation bag. | The risk is acceptable. | Instructions for use, operator training.
|  |  | R19.2 | Upload the wrong version of the program during production | V2 | Z3 | Factory technicians do not have access to multiple software versions, the software version is displayed by a number in the settings. | Based on the implemented measures, the frequency and level of risk is mitigated. | Part of the final factory check is the check the software version that is loaded on the device | V2 | Z2 | SW version check during final factory check. | The risk is acceptable. | Final check.
|  |  | R19.3 | Incorrect measurement due to damage to calibration data in memory | V2 | Z3 | The calibration data is stored in a dedicated memory, which has the validation of the stored data after switching on the system | Based on the implemented measures, the frequency and level of risk is mitigated. | Final factory check. | V2 | Z2 | Checking the measured parameters of the ventilator during the final factory check. If the memory is damaged, the device will not start. | The risk is acceptable. | Final check.
|  |  | R19.4 | Alarm failure due to a program error | V2 | Z3 | Acceptance tests were performed, which included testing the behavior of alarms | Based on the implemented measures, the frequency and level of risk is mitigated. | Final factory check, 8 hours of trial run on artificial lungs - long-term test, internal tests ... | V2 | Z3 | Stress, 8-hour tirla run / test in the laboratory including testing of alarm triggering. | The risk is acceptable. | Final check.
|  |  | R19.5 | Program memory corruption - incorrect program run | V2 | Z3 | Loading a program from the storage medium - Switch the device off and on again | Based on the implemented measures, the frequency and level of risk is mitigated. | Final factory check, 8 hours of trial run on artificial lungs - long-term test, internal tests ... | V2 | Z2 | Kontrola měřených parametrů ventilátoru během výstupní kontroly. V případě poškození paměti se přístroj nespustí. | The risk is acceptable. | Final check.
|  |  | R19.6 | Control software failure due to operating system failure | V2 | Z3 | The operating system used is optimized for embedded devices - Switching the device off and on again | Based on the implemented measures, the frequency and level of risk is mitigated. | Final factory check, 8 hours of trial run on artificial lungs - long-term test, internal tests ... | V2 | Z3 | Zátěžový, 8 hodinový test v laboratoři zahrnující testování spouštění alarmů. | The risk is acceptable. | Final check.
20. | Does the medical device have a limited shelf life? | No
21. | Are there any delayed or long-term effects? | Yes | R21.1 | Patient lung damage | V3 | Z4 | Checking the maximum pressure in the patient circuit | The risk is unacceptable, further action is needed to reduce the probability of occurrence. | Technical documentation. | V1 | Z3 | Přetlakový ventil zamezující dodávce nadlimitních tlaků do pacienta. Alarm vysokého tlaku, který zastaví ventilaci pacienta. | The risk is acceptable. | Technical documentation, Instructions for use.
|  |  | R21.3 | Release and illegibility of warning labels | V2 | Z4 | Regular inspection of the housing and placed labels | The risk is unacceptable, illegible marking of ventilation elements and places connected to the patient circuit can lead to incorrect ventilation of the patient. | Technical documentation. | V1 | Z4 | Part of regular Periodical technical safety inspections and regular maintenance is also a regular check of the legibility of labels and warnings, or their replacement if necessary. | The risk has been sufficiently reduced. | Maintenance documentation. Periodical technical safety inspection every 3 months.
|  |  | R21.4 | Damage to the patient circuit connection | V2 | Z3 | Patient circuit connection made of durable metal material. | The risk is unacceptable, regular check is necessary. | Technical documentation. | V1 | Z3 | Regular inspection by the operator of the device and during maintnance inspections. In case of damage, the device will not be used. | The risk is acceptable. | Instructions for use. Maintenance documentation. Periodical technical safety inspection every 3 months.
|  |  | R21.5 | Damage to medical gas connection hoses | V2 | Z3 | Hoses for medical gas connections made of materials according to valid standards. | The risk is unacceptable, regular check is necessary. | Technical documentation. | V1 | Z3 | Regular inspection by the operator of the device and during maintnance inspections. In case of damage, the device will not be used. | The risk is acceptable. | Instructions for use. Maintenance documentation. Periodical technical safety inspection every 3 months.
|  |  |  | CoroQuant sensor connection connection damaged. | V2 | Z3 | Regular inspection of the sensor by the operator of the device. Sensor replacement as needed - this is a disposable consumable. | The risk is unacceptable, it can lead to incorrect patient ventilation. | Instructions for use. | V1 | Z3 | Resetting the sensor by the operator according to the Instructions for use. | The risk is acceptable. | Instructions for use.
|  |  |  | CoroQuant sensor damage. | V2 | Z3 | Regular inspection of the sensor by the operator of the device. Sensor replacement as needed - this is a disposable consumable. | The risk is unacceptable, it can lead to incorrect patient ventilation. | Instructions for use. | V1 | Z3 | Resetting the sensor by the operator according to the Instructions for use. | The risk is acceptable. | Instructions for use.
|  |  |  | Damage to the ventilation circuit | V2 | Z3 | The ventilation circuit is mechanically damaged - e.g. from production. | The risk is unacceptable. | Instructions for use. | V2 | Z3 | Visual inspection of the ventilation circuit when installing it to the CoroVent ventilator, continuous checks during operation. | The risk is acceptable. | Instructions for use.
22. | What mechanical forces will the medical device exert? | Yes | R22.1
|  |  | R22.2 | Damage during storage when the product is not in use. | V2 | Z3 | After each switch-on, the device is checked by the operator before commissioning. | The risk persists during inspection by the operator only. | Instructions for use. | V2 | Z2 | After longer storage of the device without commissioning, it is recommended to check the device by a service technician. | The risk is acceptable. | Instructions for use. Maintenance documentation.
|  |  | R22.4 | Bending of supply lines during handling. | V2 | Z3 | A bend in the supply lines affecting the function of the device would result in a decrease in the working pressures of oxygen and air. The working pressures of both gases are measured by pressure sensors connected at the outlet of the inlet pressure reducing valves. The drop in working pressure is triggered by the Insufficient Oxygen Source Pressure (E2) or Insufficient Air Source Pressure (E3) alarm. | The risk is unacceptable. | Technical documentation, Instructions for use, operator training. | V2 | Z3 | Alarms of working pressures of oxygen and air. | The risk is acceptable. | Instructions for use, operator training.
23. | What determines the lifespan of a medical device? | Yes | R23.1 | Service life of used valves. | V2 | Z3 | Inspection of installed valves during assembly. The service life of the prescribed valves is 5 billion cycles = 14,000 days. | The risk is acceptable. | Datasheet of used valves, Final factory check. | V2 | Z3 | Inspection of used components - valves within the final factory check. Checking the function of the valves within the Periodical technical safety inspection. | The service life of the prescribed valves is 5 billion cycles = 14,000 days.. The risk is acceptable. | Periodical technical safety inspection every 3 months.
|  |  | R23.2 | Service life of the power supply used. | V2 | Z3 | Checking the source used during installation must be according to the proposed specification, or an alternative specified by the manufacturer. | The service life of the prescribed backup power supply is 15 years. The risk is acceptable. | Power supply and UPS certificates. | V2 | V3 | Checking the function of the backup source within the Periodical technical safety inspection | The risk is acceptable. | Periodical technical safety inspection every 3 months.
|  |  | R23.3 | Service life of used pressure sensors | V2 | Z3 | The inspection of the used pressure sensors must be according to the proposed specification, or an alternative specified by the manufacturer. | The service life of the prescribed sensors is 4 years | Certificates of used pressure sensors. | V2 | Z3 | Checking the function of the sensors whenPeriodical technical safety inspection. | The risk is acceptable. | Periodical technical safety inspection every 3 months.
|  |  | R23.4 | Operation on a backup source of electricity. | V2 | Z3 | Tested within the output control - intentional disconnection from the primary power supply. energy. Call up alarm message E12, expected full operation 20 minutes | The risk is acceptable. | Final factory check. | V2 | Z3 | Tested within the final factory check - intentional disconnection from the primary power supply. energy. Invoking alarm message E12, expected full operation 20 minutes | The risk is acceptable. | Final check.
24. | Is the medical device for one use only? | No
25. | Is it necessary to safely decommission or dispose of a medical device? | Yes | R25.1 | If the device is not disposed of as electrical waste, but as ordinary waste - environmental pollution - lead acid battery | V2 | Z2 | Check the information in the operating instructions | The risk is acceptable. | Instructions for use. |  |  | Check the information in the Instructions for use. | The risk is acceptable. | Instructions for use.
26. | Does the installation or use of a medical device require special training or special skills? | Yes | R26.1 | Failure to make a staff training | V2 | Z4 | Making personnel training to operate the CoroVent. | The risk is unacceptable. | Protokol o proškolení obsluhy. | V2 | Z3 | Zdůraznění informace o provedení školení personálu k obsluze přístroje CoroVent. | The risk is acceptable. | Records of operetor training.
27. | How will information on safe use be provided? | Yes | R27.1 | Insufficient instruction during staff training | V2 | Z4 | Training checklist | The risk is unacceptable. | Operator training report. | V2 | Z3 |  |  | Records of operetor training.
|  |  | R27.2 | Insufficient information in the instructions for use | V2 | Z4 | Instructions for use in accordance with applicable standards. | The risk is unacceptable. | Instructions for use. | V2 | Z3 | Emphasis on information on operator training. | The risk is acceptable. | Records of operetor training.
28. | Will it be necessary to create or implement new production processes? | No
29. | Does the successful use of a medical device depend decisively on human factors, such as the user interface? | Yes | R29.1 | The units of displayed quantities are different from those normally used by personnel | V2 | Z3 | Use of typical display and setting of parameters with the same quantities, which are common in practice | The risk is acceptable. | Instructions for use. | V2 | Z3 | Checking the understanding of individual indicators and measured values within the training of the operator. | The risk is acceptable. | Records of operetor training. Intructions for use.
|  |  | R29.2 | The device is perceived as a classic advanced ventilator for critical care | V2 | Z3 | Operator training, information on the typical use of the CoroVent ventilator for clinical practice. | The risk is acceptable. | Technical documentation, labels, Instructions for use, operator training. | V2 | Z3 | Technical documentation, labels, Instructions for use, operator training - all contain information on the specific use of the CoroVent ventilator. | The risk is acceptable. | Instructions for use, operator training.
29a | Can user interface elements contribute to a user error? | Yes | R29a.1 | The device is perceived as a ventilator usable for transporting a patient within a professional medical facility with only a container of compressed oxygen. | V2 | Z3 | Operating Instructions. If an attempt is made to operate the device only on a source of compressed oxygen, a low operating air pressure alarm is triggered and a warning message that the device can only be operated when connected to a source of both compressed air and oxygen. | The risk needs to be further mitigated. | Instructions for use. | V2 | Z3 | Operator training. | The risk is acceptable. | Records of operetor training.
29b | Is the medical device used in an environment where loss of attention may cause user error? | Yes | R29b.1 | User error due to non-standard appearance of the resource. | V2 | Z3 | Use of typical display and setting of parameters with the same quantities, which are common in practice. The design is chosen with maximum emphasis on the standard look | The risk needs to be further mitigated. | Instructions for use. | V2 | Z3 | Operator training. | The risk is acceptable. | Records of operetor training.
29c | Does the medical device have fasteners or accessories? | Yes | R29c.1 | Incorrect connection to the source of medical gases. | V2 | Z3 | Treated with medical quick connectors according to the standard, cannot be connected incorrectly | The risk of incorrect connection is minimized by different dimensions of quick couplings. | Technical documentation. |  |  | The correct connection of the device to the sources of medical gases is stated in the Instructions for use. | The risk has been sufficiently reduced. | Instructions for use.
29d | Does the medical device have a control interface? | Yes | R29d.1 | Difficult to operate the touch screen in protective equipment. | V2 | Z3 | The icons have been enlarged to a sufficient size to allow control of the touch dipslays in protective equipment. | The risk is acceptable when compared to the clinical benefit. | Technical documentation. | V2 | Z3 | The controls are described in the Instructions for use. | The risk is acceptable. | Instructions for use.
|  |  | R29d.2 | Difficult to read the control display due to the position of the interface on the ventilator | V2 | Z3 | The trapezoidal shape of the ventilator improves the viewing angle for the user, the readability of alarm messages has been checked | Based on the implemented measures, the frequency and level of risk is mitigated. | Technical documentation. | V2 | Z3 | Checking the display angles of the display, notices in the Instructions for use relating to the location of the ventilator and the necessary readability of the displayed data. | The risk is acceptable. | Final inspection. Instruction for use.
|  |  | R29d.3 | Difficult to operate pressure valves in protective equipment. | V2 | Z3 | Sufficiently large rotating elements. | Based on the implemented measures, the frequency and level of risk is mitigated. | Technical documentation, final factory check. | V2 | Z3 | Checking the function of adjustable parameters and controls. | The risk is acceptable. | Final inspection. Instruction for use.
|  |  | R29d.4 | Clearance in continuous control of pressure valves leading to inaccurate setting of the required parameter. | V2 | Z3 | Checking the set pressures on the manometers | Based on the implemented measures, the frequency and level of risk is mitigated. | Technical documentation - labels. | V2 | Z3 | Checking the function of adjustable parameters and controls. | The risk is acceptable. | Final inspection. Instruction for use.
|  |  | R29d.5 | Unreadable Plim and PEEP markings on pressure valves. | V2 | Z3 | Marking with non-washable labels | Based on the implemented measures, the frequency and level of risk is mitigated. | Technical documentation - labels. | V2 | Z3 | Checking the legibility of labels by the operator and during regular maintenance inspections and Periodical technical safety inspection. | Based on the implemented measures, the frequency and level of risk is reduced. | Instruction for use. Mainenance documentation. Periodical technical safety inspection every 3 months.
29e | Does the medical device display information? | Yes | R29e.1
|  |  | R29e.2 | Difficult to read the display when viewed at a certain angle of view. | V2 | Z3 | A visual verification was performed by a person familiar with the type of devices. | The risk was minimized by check. | Final factory check, Instructions for use. | V2 | Z3 | Visual inspection of the legibility of the elements, emphasis in the Instructions for use. | The risk is acceptable. | Instructions for use, operator training.
|  |  | R29e.3 | Designation of parameters, quantities and their units different from the established practices for devices of a similar type. | V2 | Z3 | Use of typical display and setting of parameters with the same quantities, which are common in practice | The risk is minimized by using values that are normal in normal practice | Instructions for use. | V2 | Z3 | Checking of the designation of parameters and values within the final factory check, operator training and alignment with the Instructions for use. | The risk is acceptable. | Instructions for use.
29f | Is the medical device controlled via a menu? | Yes | R29f.1 | Parameter and alarm settings font size too small. | V2 | Z3 | A visual verification was performed by a person familiar with the type of devices. | The risk was minimized by check. | Technical documentation, Instructions for use. | V2 | Z3 | There was a visual verification by a person familiar with the type of devices, emphasizing the fact during the training of the operator. | The risk is acceptable. | Technical documentation. Instructions for use.
|  |  | R29f.2 | the menu is confusing | V2 | Z3 | The menu was chosen to be controlled according to custom. Consulted with doctors. The data from the sensors are converted to physical units, which are displayed. | Riziko nepřehlednosti menu je minimalizováno použitím obvyklých fyzikálních veličin a nástavováním obvyklých hodnot. | Instructions for use. | V2 | Z3 | Introduction to the menu structure within the operator training, description of the menu structure in the Instructions for use. | The risk is acceptable. | Instruction for use, operetor training.
29g | Will the medical device be used by people with special needs? | No
29h | Can the user interface be used to initiate user actions? | No
30. | Does the medical device use an alarm system? | Yes
|  |  | R30.1 | False alarms. | V3 | Z3 | Inspired and expired tidal volume is calculated for each breath separately and separately. If there is a change in ventilation parameters resulting in a change in tidal volume, the Open or Leak alarms (E19, E20 and E21) may be triggered. This is a false positive trigger and no circuit disconnection has occurred. This phenomenon arises as a result of checking inspired and expired volumes and comparing them for leak monitoring purposes. False alarms are minimized by averaging the last three breath parameter values. | The risk is acceptable when compared to the benefits | Instructions for use. | V3 | Z3 | Inspired and expired tidal volume is calculated for each breath separately and indipendently. If there is a change in ventilation parameters resulting in a change in tidal volume, the Open or Leak alarms (E19, E20 and E21) may be triggered. This is a false positive trigger and no circuit disconnection has occurred. This phenomenon arises as a result of checking inspired and expired volumes and comparing them for leak monitoring purposes. | The risk is acceptable when compared to the benefits | Instructions for use.
31. | How could a medical device be intentionally misused? | Yes | R31.1 | Failure to follow the Instructions for use. | V2 | Z4 | operator training | The risk is acceptable. | Instructions for use, operator training | V2 | Z3 | Check of CoroVent operator training. | The risk is acceptable. | Operator training, Instructions for use.
|  |  | R31.2 | Deliberate overly benevolent setting of alarm limits - so that staff are not bothered by alarms | V2 | Z4 | Operator training, used only by professional staff | The risk is unacceptable, it is necessary to further mittigate the emerging risk | Records of operator training. | V2 | Z4 | Test within the final factory check - intentional triggering of alarms E10, E11, E13 | The risk is acceptable. | Instructions for use.
32. | Does the medical device store data that are crucial for patient care? | Yes | R32.1 | The device stores the set values of the ventilation parameters | V2 | Z4 | Retains the set parameters until the setting for the new patient is selected | The risk is minimized by this solution | Instructions for use. | V2 | Z3 | Test within the final factory check - display of an information window with the appropriate query | Risk minimized. | Instructions for use.
|  |  | R32.2 | The device stores the values of the measured parameters | V2 | Z2 | It stores the measured parameters in memory without the patient's personal information | There is no risk | Technical documentation. | V2 | Z2 | It stores the measured parameters in memory without the patient's personal information | There is no risk | Technická dokumentace
33. | Is the medical device intended to be mobile or portable? | Yes | R33.1 | Falling of the ventilator to the ground during transport, relocation, handling | V2 | Z2 | Trapezoidal shape with sufficient base | The risk is minimized by this solution. | Instructions for use, operator training. | V2 | Z2 | Trapezoidal shape with sufficient base | The risk is minimized by this solution | Instructions for use, operator training.
|  |  | R33.2 | Difficult handling of the ventilator due to the dimensions and the absence of handles | V2 | Z2 | Weight less than 20 kg can be transferred by holding the lower edge of the ventilator | There is no risk of problematic manipulation. | Technical documentation, Instructions for use. |  |  | Weight less than 20 kg, can be transported by holding the lower edge of the ventilator | There is no risk of problematic manipulation | Instructions for use, operator training.
|  |  | R33.3 | Due to the shape of the ventilator housing, the risk of "tipping over" | V2 | Z2 | Trapezoidal shape with sufficient base | The risk was sufficiently mitigated by the shape of the ventilator. | Technická dokumentace |  |  | Trapezoidal shape with sufficient base | The risk was sufficiently reduced by the shape of the ventilator | Instructions for use, operator training.
34. | Does the use of a medical device depend on the basic function? | Yes | R34.1 | Connecting the ventilator only to a source of compressed oxygen without a source of compressed air that drives the ventilating expiratory branch of the ventilator | V2 | Z4 | Alarm messages E3, E99 | The risk is unacceptable, it is necessary to further mitigate the emerging risk. | Instructions for use. | V2 | Z3 | Instruction for us, operator training. | The risk has been sufficiently reduced. | Instructions for use, operator training.
35. | Can a medical device cause biological or chemical hazards? | Yes | R35.1 | Unwanted microorganisms can multiply in the ventilation circuit | V3 | Z4 | Regular replacement of ventilation circuits | The risk is unacceptable. | Instructions for use. | V3 | Z3 | Reduction of the time of use of the ventilation circuit and HME filter - interval according to the Instructions for use. | The risk has been sufficiently reduced. | Instructions for use, operator training.
36 | Unclassified risks | Yes | R35 | Ensuring the opening of the oxygen reducing valve | V2 | Z4 | Incoming inspection before assembly. | The risk is unacceptable, it is necessary to further mitigate the emerging risk. | Assembly documentation. | V2 | V3 | Alarm messages E2, E99 | The risk has been sufficiently reduced. | Instructions for use.
|  |  | R36 | replacement of hoses when measuring pressures and flows in the patient circuit | V2 | Z3 | Color differentiation | The risk is unacceptable, it is necessary to further mitigate the emerging risk. | Instructions for use. | V2 | Z3 | Operator training to reset the flow sensor and to check the direction of gas flow vs, monitored quantity. | The risk has been sufficiently reduced. | Instructions for use, operator training.
|  |  |  | Incorrectly set output pressure value on the O2 pressure reducing valve | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | The station of setting the outlet pressure of the O2 pressure reducing valve is incorporated into the assembly line. Description in the Assembly Manual. Monitored parameter during final factory check of the device. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Malfunction of the 2/2 Festo valve | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | The correct function of the 2/2 Festo valve is subject to final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Leakage of compliance containers | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | A ductility leak test station has been integrated into the assembly line. Description in the Assembly Manual. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Incorrectly set outlet pressure value on the compressed air pressure reducing valve | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Test station for setting the outlet pressure of the compressed air pressure reducing valve is integrated into the assembly line. Description in the Assembly Manual. Monitored parameter during final factory check of the device | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Expiration valve malfunction | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | An expiration valve test station is integrated into the assembly line. Description in the Assembly Manual. Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Incorrectly set overpressure relief valve | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | A test station for overpressure safety valve is integrated into the assembly line. Description in the Assembly Instructions. Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Leakage of the entire pneumatic circuit of the device | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Incorrect function of the backup power supply. | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Verification of the functionality of the backup power source during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Leakage of the connecting hose of the compressed air. | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Test position for checking the tightness of the connecting hose of the compressed air is integrated into the assembly line. Description in the Assembly Instructions. Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Leakage of the connection hose of the O2. | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Test position for checking the tightness of the connecting hose of the O2 is integrated into the assembly line. Description in the Assembly Instructions. Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Mechanical impurities inside pneumatic elements | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Test position for blowing all pneumatic elements with compressed air is integrated into the assembly line. Description in the Assembly Instructions. Monitored parameter during final factory check. | The risk was sufficiently mitigated. | Assembly Manual
|  |  |  | Paw pressure gauge malfunction | V2 | Z3 | Final factory check of the product. | The risk is unacceptable, it is necessary to further mitigate the risk. | Assembly Manual | V2 | Z3 | Modification of the design of the Paw manometer, making it impossible to plug the hole for equalizing the pressure of the manometer. Subject to final factory check. | The risk was sufficiently mitigated. | Assembly Manual
Abreviations:
HME | Humidifier Moisture Exchanger
VDO
ZIS
EMC | Electro Magntic Compatibility
EZU | "Elektrotechnicky zkusebni ustav" - Elektrotechnical test institut